EEG Signal Modeling in LTspice

Summary of the behavioral signal, differential model, common-mode interference, and electrode mismatch study

Objective. Develop a controllable EEG stimulus in LTspice for future CMOS analog front-end evaluation. The model was built progressively from a clean multi-band signal to a differential input with 50 Hz common-mode interference and unequal electrode impedances.

# 1. Scope and simulation setup

Tool: LTspice behavioral voltage sources (bv).

Transient analysis: .tran 0 10 0 100u.

Nominal input common-mode voltage: 0.9 V.

Main EEG range obtained: approximately −46 µV to +54 µV.

The model is deterministic and intended as an engineering test stimulus, not as a clinical EEG recording.

.tran 0 10 0 100u

# 2. Stage 1 — Clean composite EEG signal

The first source was constructed as the sum of four sinusoidal components representing the main EEG frequency bands:

| Component | Frequency | Amplitude |
| --- | --- | --- |
| Delta | 2 Hz | 20 µV peak |
| Theta | 6 Hz | 10 µV peak |
| Alpha | 10 Hz | 30 µV peak |
| Beta | 20 Hz | 5 µV peak |

V=20u*sin(2*pi*2*time)+10u*sin(2*pi*6*time+0.4)+30u*sin(2*pi*10*time+1)+5u*sin(2*pi*20*time+0.2)

Verification: The individual delta, theta, alpha, and beta sources were plotted separately. Their sum matched V(EEG). The error trace below stayed essentially at zero:

V(eeg)-V(delta)-V(theta)-V(alpha)-V(beta)

Observation: Because all selected frequencies are integer multiples of 2 Hz, the composite waveform repeats every 0.5 s. This is acceptable for the initial controlled model.

# 3. Stage 2 — Differential EEG representation

The single-ended EEG waveform was converted into two differential input nodes around a 0.9 V common mode:

V(INP_IDEAL) = V(VCM) + 0.5*V(EEG)
V(INN_IDEAL) = V(VCM) - 0.5*V(EEG)

Each input contains half of the EEG excursion.

The full EEG is recovered by subtraction: V(INP_IDEAL)−V(INN_IDEAL)=V(EEG).

The average of both inputs remains at 0.9 V: [V(INP_IDEAL)+V(INN_IDEAL)]/2=0.9 V.

Important plotting note: The individual inputs appear flat when shown on a volt-scale axis because the useful variation is only tens of microvolts. The differential trace must be plotted separately or the vertical axis rescaled to microvolts.

# 4. Stage 3 — 50 Hz common-mode interference

A large 50 Hz common-mode interference was added to both inputs:

V(VCM_NOISY)=0.9+100m*sin(2*pi*50*time)

V(INP_IDEAL_NOISY)=V(VCM_NOISY)+0.5*V(EEG)
V(INN_IDEAL_NOISY)=V(VCM_NOISY)-0.5*V(EEG)

VCM_NOISY swings from about 0.8 V to 1.0 V.

INP_IDEAL_NOISY and INN_IDEAL_NOISY visually overlap because the 100 mV common-mode term dominates the microvolt EEG difference.

With ideal and perfectly matched paths, the 50 Hz term cancels completely in the differential result.

V(INP_IDEAL_NOISY)-V(INN_IDEAL_NOISY)=V(EEG)

Common-mode verification: The following average reproduces V(VCM_NOISY):

(V(INP_IDEAL_NOISY)+V(INN_IDEAL_NOISY))/2

# 5. Stage 4 — Unequal electrode impedances

To demonstrate common-mode-to-differential conversion, unequal source resistances and finite input resistance were added:

INP_IDEAL_NOISY — 50k — INP — 100Meg — VCM_BIAS
INN_IDEAL_NOISY — 70k — INN — 100Meg — VCM_BIAS

VCM_BIAS is a clean 0.9 V DC reference.

The 100 MΩ resistors represent a simplified finite input resistance and provide a small current path.

Without this finite input resistance, no current would flow through the 50 kΩ and 70 kΩ source resistances, so their mismatch would have no effect.

Because the divider ratios are different, equal common-mode interference reaches INP and INN with slightly different amplitudes.

Approximate divider gains:

A_P = 100M/(100M+50k)
A_N = 100M/(100M+70k)

Expected residual: For a 100 mV common-mode sine wave, the mismatch produces approximately 20 µV peak of differential 50 Hz error.

The error was isolated using:

(V(INP)-V(INN))-V(EEG)

The resulting waveform was a clean 50 Hz sinusoid with approximately ±20 µV amplitude.

The corrupted signal at the CMOS input was therefore: V(INP)−V(INN)=V(EEG)+Verror,50Hz.

This demonstrates that electrode/source impedance mismatch can degrade common-mode rejection even before amplifier non-idealities are included.

# 6. Completed model summary

| Stage | Model element | Key result |
| --- | --- | --- |
| 1 | Four-band clean EEG source | Controlled waveform of tens of microvolts |
| 2 | Differential conversion around 0.9 V | Differential subtraction reproduces the EEG |
| 3 | 100 mV, 50 Hz common-mode interference | Perfect cancellation with ideal matched paths |
| 4 | 50 kΩ/70 kΩ source mismatch + 100 MΩ input resistance | About 20 µV peak residual 50 Hz differential error |

# 7. Progression — frequency-dependent electrode model

The next stage is to replace each fixed electrode resistor with a first-order electrode–skin interface model:

Z_electrode = Rs + (Rct || Cdl)

Rs models series resistance of tissue/contact.

Rct models charge-transfer or leakage resistance.

Cdl models the double-layer capacitance at the electrode interface.

The resulting impedance becomes frequency dependent, so both attenuation and phase mismatch vary with frequency.

Planned illustrative values:

| Path | Rs | Rct | Cdl |
| --- | --- | --- | --- |
| Positive electrode | 10 kΩ | 500 kΩ | 100 nF |
| Negative electrode | 15 kΩ | 700 kΩ | 68 nF |

Planned connections: INP_IDEAL_NOISY → Rs,P → [Rct,P ∥ Cdl,P] → INP, and similarly for the negative path. The existing 100 MΩ input resistors and 0.9 V VCM_BIAS remain in place.